Tender Intake Report (Pre-Bid)

Tender Lab Automation Line - Project description — 13.12.2024 (Insel Gruppe AG, Berne, Switzerland)

Generated: 2026-02-01 19:12

1. Executive Summary

Scope: turnkey Lab Automation Line replacement including full responsibility for system integration and third-party coordination.

Critical constraints: replacement during ongoing operations; fixed sample collection area; strict space boundaries across INO F and INO E; mandatory biobank connection.

Key milestones identified: contract in 2Q2025, middleware/training from 3Q2025, installation start January 2026.

Pre-bid recommendation: GO with Mitigation (complexity 61/100). Focus on early feasibility checks and integration responsibilities.

2. Key Deadlines & Milestones

| Milestone | When | Evidence |
| --- | --- | --- |
| Contract signature | 2Q2025 | p.5 |
| Middleware configuration & training | from 3Q2025 | p.5 |
| Installation start | January 2026 | p.5 |

3. Requirements & Constraints (Extracted)

Project scope & responsibility

Bidder responsible for entire system; delivers a turnkey system; third-party providers contracted directly by bidder. (p.2)

Project description is evaluated as part of bids. (p.2)

Required content in project description

Layout of the Lab Automation Line (p.2)

Project scheduling (planning/project planning, system delivery, installation, testing & commissioning) (p.2)

Parts list of the components offered (p.2)

The provider shall examine the feasibility of a layout exclusively on the INO F floor as a variant. (p.2)

Operational / brownfield constraints

The replacement of the existing Lab Automation Line must be realized during ongoing operations (p.3)

System setup and commissioning coordinated with overall project management of hospital/clinic. (p.3)

Sample collection area remains unchanged. (p.3)

Facility / space constraints

The space in INO F shown in the enclosed planning documents is to be used (primarily the yellow area in the Core Lab is available; if more space is required, the green and secondarily the blue area should be used as a priority) (p.3)

Use INO E space shown in planning documents. (p.3)

Biobank moved to INO E; connection of biobank to Lab Automation Line is a requirement. (p.3)

Door insertion heights: INO F 2.20 m; INO E 2.10 m (width 2.00 m). (p.8-9)

Ceiling height constraints in INO E (>=2.40 m in highlighted areas; 2.20–2.40 m elsewhere). (p.9)

Potential ceiling breakthrough area for vertical connection between floors. (p.10-11)

Analyzer connectivity & layout rules

Analytics per tender documentation must be connected to Lab Automation Line; reserve space for additional CN-6000 and Alinity i. (p.3)

Group devices along the track by specialist areas (work economy). (p.3)

Existing analytics not connected to track must be relocated to the perimeter. (p.3)

If not possible to fully connect all analytics due to space, apply prioritization list (Cobas Pro, CN-6000, Liaison XL, Alinity i, Phadia 250, plus reserves). (p.4)

Provide a validation workstation for each analysis. (p.4)

Schedule / milestones

Signing of contract 2Q2025 (p.5)

Configuration of middleware, training, etc. from 3Q2025 (p.5)

Start of installation of the Lab Automation Line January 2026 (p.5)

Provided drawings / supplements

Surveyed plan of the laboratory perimeter (the automation of the biobank is not shown) File name: Laborperimeter_Geschoss INO F (p.6)

Floor plans of floors INO E and INO F in DWG format (ATTENTION: the third column is not shown in GR INO F / the washbasin at column G/5 does not exist) File names: Floor plan - 056E & Floor plan - 056F (p.6)

Master plan of the automation of the biobank in DWG format File name: Bern-Inselspital-Biobank connection (p.6)

Schematic overview plan of the current Lab Automation Line including the automated biobank (in the appendix of this document, page 7) (p.6)

Overview plan of the areas available in INO F and INO E (in the appendix to this document, pages 8 & 9) (p.6)

Overview plans with the intersection area in INO F and INO E for the vertical connection of the two floors (in the appendix of this document, pages 10 & 11) (p.6)

4. Deliverables to Prepare (Pre-Bid)

Pre-bid feasibility assessment (layout variant exclusively on INO F floor).

High-level layout(s) for INO F / INO E including biobank connection concept.

Project schedule (planning → delivery → installation → testing & commissioning).

Parts list / BOM of offered components (high level).

Assumptions & deviations log (justify any deviation from stated framework conditions).

Interface & responsibility matrix for third-party providers (turnkey scope).

Site survey / drawing request list (DWGs: INO E/F plans, perimeter, biobank master plan).

5. Risk Register (Top)

| ID | Risk | Category | Prob (1-5) | Impact (1-5) | Score | Evidence |
| --- | --- | --- | --- | --- | --- | --- |
| R1 | Replacement during ongoing operations (brownfield). Potential impact on continuity, phased cutover complexity. | Operational / Timeline | 4 | 5 | 20 | p.3 |
| R2 | Turnkey responsibility incl. direct contracting of third parties: commercial/legal exposure, integration risk. | Commercial / Legal | 3 | 4 | 12 | p.2 |
| R3 | Biobank relocation to INO E and mandatory automated connection to Lab Automation Line (multi-floor integration). | Technical / Facility | 4 | 4 | 16 | p.3, p.10-11 |
| R4 | Space constraints and strict area prioritization in INO F; may force partial connectivity and compromises. | Facility / Layout | 4 | 4 | 16 | p.3, p.8 |
| R5 | Requirement to connect multiple analyzer brands + provide validation workstation per analysis; footprint and workflow complexity. | Technical / Workflow | 3 | 4 | 12 | p.3-4 |

6. Go / No-Go Recommendation

Overall complexity score: 61 / 100

Recommendation: GO with Mitigation

Key drivers are brownfield replacement during operations, mandatory biobank connection across floors, and tight space constraints. Proceed only if early feasibility (INO F-only variant + biobank link concept) is confirmed via drawings/site survey and responsibilities for third parties are contractually manageable.